## Air Density
Air Density
| as in Cal. Guide |  |  | as in PP-slide
ta | 20 |  |  | 20
pa | 1014 |  |  | 1014
hr | 50 |  |  | 50
k1 | 0.00034844 |  |  | 0.34844
k2 | -2.52e-06 |  |  | -0.00252
k3 | 2.0582e-05 |  |  | 0.020582
ra | 0.0012002 | g/mL |  | 1.2002 | g/L
ra | 1.2002 | g/L
ra,ref | 1.2002 |  | at 20°C / 1014 hPa / 50% hr
Dra | -0.003232%

## Water Density
Density of bi-distilled water
| a1 = | -3.983 | A = | 999.8531
| a2 = | 301.797 | B = | 0.063269
| a3 = | 522528.9 | C = | -0.0085238
| a4 = | 69.3488 | D = | 6.9432e-05
| a5 = | 999.975 | E = | -3.8212e-07
| t / ºC | r / g·L-1 | r / g·L-1 | r(t)-r(20) / kg/m3 | r(t)-r(20)
|  | new (2001) | old
| 0 | 999.8428 | 999.8531 | 1.6498 | 0.165%
| 1 | 999.9017 | 999.9079 | 1.7046 | 0.1705%
| 2 | 999.9429 | 999.9461 | 1.7428 | 0.1743%
| 3 | 999.9672 | 999.968 | 1.7648 | 0.1765%
| 4 | 999.9749 | 999.9741 | 1.7709 | 0.1771%
| 5 | 999.9668 | 999.9648 | 1.7615 | 0.1762%
| 6 | 999.9431 | 999.9403 | 1.7371 | 0.1737%
| 7 | 999.9045 | 999.9012 | 1.6979 | 0.1698%
| 8 | 999.8513 | 999.8477 | 1.6444 | 0.1645%
| 9 | 999.7839 | 999.7802 | 1.5769 | 0.1577%
| 10 | 999.7027 | 999.699 | 1.4957 | 0.1496%
| 11 | 999.6081 | 999.6045 | 1.4012 | 0.1401%
| 12 | 999.5005 | 999.4969 | 1.2937 | 0.1294%
| 13 | 999.3801 | 999.3767 | 1.1734 | 0.1174%
| 14 | 999.2474 | 999.244 | 1.0408 | 0.1041%
| 15 | 999.1026 | 999.0992 | 0.89599 | 0.08961%
| 16 | 998.9459 | 998.9426 | 0.73939 | 0.07395%
| 17 | 998.7778 | 998.7745 | 0.57122 | 0.05713%
| 18 | 998.5984 | 998.595 | 0.39177 | 0.03918%
| 19 | 998.4079 | 998.4045 | 0.20128 | 0.02013%
| 20 | 998.2067 | 998.2033 | 0 | 0%
| 21 | 997.995 | 997.9914 | -0.21183 | -0.02119%
| 22 | 997.773 | 997.7693 | -0.43398 | -0.0434%
| 23 | 997.5408 | 997.537 | -0.66623 | -0.06663%
| 24 | 997.2988 | 997.2949 | -0.90838 | -0.09085%
| 25 | 997.047 | 997.043 | -1.1602 | -0.116%
| 26 | 996.7857 | 996.7817 | -1.4216 | -0.1422%
| 27 | 996.5151 | 996.511 | -1.6922 | -0.1692%
| 28 | 996.2353 | 996.2312 | -1.972 | -0.1972%
| 29 | 995.9465 | 995.9425 | -2.2608 | -0.2261%
| 30 | 995.6488 | 995.6449 | -2.5584 | -0.2559%
| 31 | 995.3424 | 995.3386 | -2.8647 | -0.2865%
| 32 | 995.0275 | 995.0238 | -3.1795 | -0.318%
| 33 | 994.7041 | 994.7005 | -3.5027 | -0.3503%
| 34 | 994.3724 | 994.369 | -3.8342 | -0.3835%
| 35 | 994.0326 | 994.0293 | -4.1739 | -0.4175%
| 36 | 993.6847 | 993.6815 | -4.5217 | -0.4522%
| 37 | 993.329 | 993.3257 | -4.8775 | -0.4878%
| 38 | 992.9654 | 992.962 | -5.2412 | -0.5242%
| 39 | 992.5941 | 992.5905 | -5.6128 | -0.5614%
| 40 | 992.2152 | 992.2112 | -5.9921 | -0.5993%

## Template
Calibración gravimétrica de una probeta
| Quantity | Value | Uncertainty | Distribution | d.f. | Standard Uncertainty | Sensitivity Coefficient | Uncertainty Contribution | "Index" (Variance)
i |  | xi |  |  | vi | u(xi) | ci | ci*ui | (ci*ui/uc)2
1 | Weighing result: W
1c | Repeatibility: Ŵ, dWrep
1a | Calibration: DWcal, dWcal
1b | Resolution: dWres
2 | Air buoyancy: Cbuoy,dCbuoy
3 | Water density: r
Volume: V

## U-budget (4)
Calibración gravimétrica de una probeta
| Quantity | Value | Uncertainty | Distribution | Standard Uncertainty | Sensitivity Coefficient | Uncertainty Contribution | "Index" (Variance)
i |  | xi |  |  | u(xi) | ci | ci*ui | (ci*ui/uc)2
1 | Weighing result: W | 1992.8
1c | Repeatibility: Ŵ, dWrep | 1993 | 1 | normal | 1
1a | Calibration: DWcal, dWcal | -0.2 | 1.2 | normal, k=2 | 0.6
1b | Resolution: dWres | 0 | 1 | rectangular | 0.28868
2 | Air buoyancy: Cbuoy,dCbuoy | 2.4 | 0.1 | rectangular | 0.057735
3 | Water density: r | 998.2 | 0.4 | rectangular | 0.23094
| Volume: V | 1.9988

## U-budget complete
Calibración gravimétrica de una probeta
| Quantity | Value | Uncertainty | Distribution | Standard Uncertainty | Sensitivity Coefficient | Uncertainty Contribution | "Index" (Variance) | Degrees of freedom | (ui/uc)4/vi
i |  | xi |  |  | u(xi) | ci | ci*ui | (ci*ui/uc)2
1 | Weighing result: W | 1992.8
1c | Repeatibility: Ŵ, dWrep | 1993 | 1 | normal | 1 | 0.0010018 | 0.0010018 | 60.25% | 4 | 0.090757
1a | Calibration: DWcal, dWcal | -0.2 | 1.2 | normal, k=2 | 0.6 | 0.0010018 | 0.00060108 | 21.69% | 10000 | 4.7048e-06
1b | Resolution: dWres | 0 | 1 | rectangular | 0.28868 | 0.0010018 | 0.0002892 | 5.021% | 10000 | 2.521e-07
2 | Air buoyancy: Cbuoy,dCbuoy | 2.4 | 0.1 | rectangular | 0.057735 | 0.0010018 | 5.7839e-05 | 0.2008% | 10000 | 4.0336e-10
3 | Water density: r | 998.2 | 0.4 | rectangular | 0.23094 | 0.0020022 | 0.00046239 | 12.84% | 10000 | 1.6476e-06
| Volume: V | 1.9988 |  |  |  | uc = | 0.0012906 |  | 11.0176 | 0.090764
|  |  | not considering degrees of freedom |  |  | U95,45 = | 0.0025812
[6 empty] | t95,45 = | 2.2549
|  |  | considering degrees of freedom |  |  | U95,45 = | 0.0029102
[6 empty] | Vlow = | 1.9959
[6 empty] | Vhigh = | 2.0017
[6 empty] | Vlow = | 1.9962
[6 empty] | Vhigh = | 2.0014