Final Report – Draft B

Bilateral comparison

Of the turbine gas meter G1000

BILATERAL comparison

EURAMET Project No. 1007

Dr.Vahit Çiftçi – Hakan Kaykısızlı – Alper Elkatmış

(National Metrology Institute of Turkey)

Mijndert P. Van Der Beek - Rob Zwaan

(National Metrology Institute of Netherlands)

January 2010

Contents

- Introduction

- The instrument

- Calibration procedure

- Equipment and obtained results

  - Netherlands

  - Turkey

- Consistency of calibrations

  - Consistency of UME

- Bi-lateral Results

- Summary and conclusion of EUROMET 1007 bilateral comparison

- Comparison of EUROMET 862 and EUROMET 1007 measurement results

- Summary and conclusion of comparison between EUROMET 862  and EUROMET 1007  comparison results

- Introduction

The project EUROMET no.862 for the comparison of the axial turbine gas meter G650 concluded in October 2006.  Each country performed the calibration of the turbine gas meter G650 with air in the barometric pressure. The range of flow rates was from 100 m3/h to 1000 m3/h.  The participating laboratories used their usual calibration procedure. When the final draft of the comparison published, the results showed that something was wrong with the calibration of Turkey (UME).

Turkey (UME) wanted to investigate the reasons for this unexpected deviation form the mean of the other countries.  Turkey’s (UME) gas measurement reference system was a combination of reference gas flow meters calibrated by the Netherlands (NMi), so a bilateral comparison was offered to check if something was wrong with the calibrations of UME reference standards. The aim of the bilateral comparison is to compare the performance of the calibrations of the gas meter in UME and NMI gas flow laboratories in order to investigate the need for a back-dated correction of UME reference K-factor according to the
results of Euromet project 862. An axial turbine gas meter G1000 was used as the decisive meter to compare low-pressure test facilities over the flow range (80 ÷ 1600) m3/h. It was decided to conduct this bilateral comparison during the Euramet Flow meeting held in Istanbul in 2007.

Table 1 – Time schedule

| Country | Laboratory | Address of the Place of calibration | Date of calibration | Responsible |
| --- | --- | --- | --- | --- |
| Turkey (ATA-CARNET) | TUBITAK – UME National Metrology Institute of Turkey | TUBITAK-UME Anibal Cad. MAM Kampusu PK54 - 41470 Gebze-Kocaeli / TURKEY | 01.4.-01.5. 2007 | Vahit Ciftci |
| Netherlands | NMi VSL Nederlands Meetinstituut Van Swinden Laboratorium | NMi VSL Flow Thijsseweg 11, 2629JA Delft, The Netherlands | 01.6.-01.10. 2007 | Mijndert P. van der Beek |
| Turkey (ATA-CARNET) | TUBITAK – UME National Metrology Institute of Turkey | TUBITAK-UME Anibal Cad. MAM Kampusu PK54 - 41470 Gebze-Kocaeli / TURKEY | 01.10.-01.12. 2007 | Vahit Ciftci |

- The instrument

The axial turbine gas meter (Fig. 1) was used for the comparison. The description of this meter is given below.

| Manufacturer: VemmTec GERMANY |  |
| --- | --- |
| Type: G1000/DN200/200 | ANSI300/PN40 Pmax : 60 bar |
| Size: G1000 | Inside diameter: DN 200 |
| Serial number: 031271 | Pulse number HF1: 2743,08 imp/m3 |
| Qmin : 80 m3/h | HF3 : 144,373 imp/m3 |
| Weight: 100 kg | Qmax : 1600 m3/h |

Figure 1 – Axial turbine gas meter, type IGTM

The dimensions of the meter are given in table 2 and in figure 2.

Table 2  - Dimensions of  the meter

| L | D | E | H | A | B | C |
| --- | --- | --- | --- | --- | --- | --- |
| 600 mm | 375 mm | 275 mm | 270 mm | 240 mm | 278 mm | 195 mm |

Figure 2 - Dimensions of the meter

- Calibration procedure

- The turbine gas meter is tested in horizontal position by air.

- For the test it was necessary to use the upstream straightening pipe with a length of at least 5xDN.

- For the test it was necessary to use the downstream straightening pipe that is long with a length of at least 3xDN.

- The reference pressure from the turbine gas meter was measured from the output “pr”.

- The reference temperature from the turbine gas meter was measured at the distance 2 x DN downstream of the turbine gas meter.

- For the test it was necessary to use both pulse outputs of the test meter, HF1 and HF3.

- The test should be performed in a laboratory where the temperature is from 19.5°C to 23.5°C.  The upstream pressure of the meter should be near atmospheric pressure.

- The oil pump of the turbine gas meter is not to be used.

- Before the beginning of the test, the gas meter worked 20 minutes at nominal flow rate Q=800 m3/h.

- The turbine gas meter was tested at 6 flow rates: 1600 m3/h, 1100 m3/h, 650 m3/h, 400 m3/h, 160 m3/h and 80 m3/h.

- The test at one flow rate should be repeated at least 3 times and then the means of values in the table 3 and table 4 have to be calculated.  The flow rate has to be in the interval ± 3% of the required value.

- A single test in one flow rate has to take more then 1 minute. Beforehand, the flow rate has to be accurately stabilised.

- Each participant recorded the results in the form of table 3 given below for HF1 and HF3 output results of the test meter (decisive meter).

UME will use the G1000 type master meter (serial number 031265) for calibrating the test meter. This master meter has two HF pulse outputs; 145,259 and 2759,92 imp/m3 respectively. UME will use both outputs of the master meter for this bilateral comparison.  Because the K-factor of this master meter was a point of discussion in Euromet 862 project, analysis of the results of new NMI certificate for the master meter will also be mentioned in the final report. So in the new NMI certificate, results of both HF pulse outputs of the meter with serial number 031265 were given to investigate the problem with the UME results for Euromet 862.

Table 3 - Required table of results

Table3

| Test Flow rates | Reference standard, flow rate at working Condition | Test meter flow rate at working Condition by HF1 | Absolute pressure in the meter | Temperature in the meter | Reference standard, flow rate at Standard Condition | Test meter flow rate at Standard Condition by HF1 | Error for the Standard Conditions for HF1 | Uncertainty of the error U (k=2) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (m3/h) | (m3/h) | (m3/h) | (Pa) | (°C) | (Sm3/h) | (Sm3/h) | (%) | (%) |
| 1600 |  |  |  |  |  |  |  |  |
| 1100 |  |  |  |  |  |  |  |  |
| 650 |  |  |  |  |  |  |  |  |
| 400 |  |  |  |  |  |  |  |  |
| 160 |  |  |  |  |  |  |  |  |
| 80 |  |  |  |  |  |  |  |  |

Table 3 - Required table of results

Table4

| Test Flow rates | Reference standard, flow rate at working Condition | Test meter flow rate at working Condition by HF1 | Absolute pressure in the meter | Temperature in the meter | Reference standard, flow rate at Standard Condition | Test meter flow rate at Standard Condition by HF1 | Error for the Standard Conditions for HF1 | Uncertainty of the error U (k=2) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (m3/h) | (m3/h) | (m3/h) | (Pa) | (°C) | (Sm3/h) | (Sm3/h) | (%) | (%) |
| 1600 |  |  |  |  |  |  |  |  |
| 1100 |  |  |  |  |  |  |  |  |
| 650 |  |  |  |  |  |  |  |  |
| 400 |  |  |  |  |  |  |  |  |
| 160 |  |  |  |  |  |  |  |  |
| 80 |  |  |  |  |  |  |  |  |

Error of the meter is value which shows the relationship in percentage terms of the difference between the volume indicated by the meter and the volume which has actually flowed through the meter, to the later value.

E=Vi−VcVc.100                (%)                                           [1]

Where     E is the error of the meter

Vi   is the indicated volume by the meter (m3)

Vc   is the real volume which has actually flowed through the meter (m3)

Error of the meter is value which shows the relationship in percentage terms of the difference between the volume indicated by the meter and the volume which has actually flowed through the meter, to the later value.

E=Vi−VcVc.100                (%)                                           [1]

Where     	E is the error of the meter

Vi   is the indicated volume by the meter (m3)

Vc   is the real volume which has actually flowed through the meter (m3)

- Equipment and obtained results

  - Netherlands

The turbine meter used for this bilateral comparison has been calibrated on the ‘Big Installation’ of NMi VSL, located in our lab in Delft, Thijsseweg 11.The installation is made up of two axial blowers, one for low flow rates (up to 4000 m3/h) and one for high flow rates (up to 15000 m3/h), a regulated heat exchanger after the blowers, data acquisition system, piping, pressure and temperature sensors and five reference meters. The reference meters are described in table 4: Nmi Gas Flow rate Measurement system’s total uncertainty is about 0.15 % (for k=2) at this range.

Table 4: Reference meters in big installation in NMi

Results of Netherlands for UME Transfer meter with serial number 031271:

Table5

| Test Flow rates | Reference flowrate [m3/h] | Indicated flow rate [m3/h] | Pabs MUT [Pa] | Tabs MUT [°C] | Reference flowrate @ Standard Conditions [sm3/h] | Indicated flow rate @ Standard Conditions [sm3/h] | Error HF1 [%] | Uncertainty [%] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1600 | 1605,0 | 1603,1 | *** | *** | *** | *** | -0,12 | 0,19 |
| 1100 | 1103,0 | 1102,0 | *** | *** | *** | *** | -0,09 | 0,19 |
| 650 | 645,3 | 644,0 | *** | *** | *** | *** | -0,20 | 0,19 |
| 400 | 394,9 | 394,0 | *** | *** | *** | *** | -0,22 | 0,19 |
| 160 | 158,1 | 157,0 | *** | *** | *** | *** | -0,68 | 0,20 |
| 80 | 78,9 | 77,3 | *** | *** | *** | *** | -2,09 | 0,20 |

  - Turkey

## UME - Medium Flow Gas Measurement System Description

As seen from figure 13, UME Medium Gas Flow Rate Measurement System was designed to operate in the range of 5-4800 m3/h for gas flow meter calibrations. It consists of one flow computer, two suction fans with frequency controller, five reference flow meters with a different measuring capacity and upstream and downstream of each flow meter has pressure-actuated ball valves. Time measurement, pulse counts, temperature and pressure at upstream of each flow meters are recorded by the UME flow computer.

Flow meters are traceable to NMi and UME. UME Medium Gas Flow rate Measurement system’s total uncertainty is about 0.18 % (for k=2).

Figure 13   – Test bench in UME

Date: 30.04.2007, First Comparison result for TS- 031271

Each period for each measurement, data was taken from the output signals of Transfer Standard’s HF1 and HF3 channels and results were recorded; it has been seen from the results that there was no significant difference between two signals data.

Table 7

| Test Flow rates | Reference standard, flow rate at working Condition | Test meter flow rate at working Condition by HF1 | Absolute pressure in the meter | Temperature in the meter | Reference standard, flow rate at Standard Condition | Test meter flow rate at Standard Condition by HF1 | Error for the Standard Conditions for HF1 | Uncertainty of the error U (k=2) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (m3/h) | (m3/h) | (m3/h) | (Pa) | (°C) | (Sm3/h) | (Sm3/h) | (%) | (%) |
| 80 | 81,298 | 78,162 | 995,63 | 20,00 | 79,861 | 76,803 | -3,830 | 0,20 |
| 160 | 161,651 | 159,214 | 995,60 | 20,17 | 158,671 | 156,350 | -1,463 | 0,20 |
| 400 | 403,349 | 401,268 | 995,42 | 20,33 | 395,427 | 393,763 | -0,421 | 0,20 |
| 650 | 653,611 | 650,433 | 995,27 | 20,41 | 639,857 | 637,999 | -0,290 | 0,20 |
| 1100 | 1102,637 | 1095,025 | 994,71 | 20,14 | 1076,973 | 1074,476 | -0,232 | 0,20 |
| 1600 | 1592,771 | 1575,480 | 993,79 | 20,41 | 1545,272 | 1543,064 | -0,143 | 0,20 |

Date: 30.11.2007, Second Comparison result for TS-031271

Table 8

| Test Flow rates | Reference standard, flow rate at working Condition | Test meter flow rate at working Condition by HF1 | Absolute pressure in the meter | Temperature in the meter | Reference standard, flow rate at Standard Condition | Test meter flow rate at Standard Condition by HF1 | Error for the Standard Conditions for HF1 | Uncertainty of the error U (k=2) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (m3/h) | (m3/h)4 | (m3/h) | (Pa) | (°C) | (Sm3/h)5 | (Sm3/h) | (%) | (%) |
| 80 | 85,350 | 83,873 | 1002,57 | 20,70 | 84,2305 | 82,791 | -1,709 | 0,20 |
| 160 | 157,850 | 156,672 | 1002,68 | 21,38 | 155,4326 | 154,314 | -0,720 | 0,20 |
| 400 | 412,045 | 410,891 | 1000,37 | 20,26 | 405,8387 | 405,312 | -0,130 | 0,20 |
| 650 | 653,818 | 651,915 | 999,09 | 20,42 | 642,6812 | 641,878 | -0,125 | 0,20 |
| 1100 | 1097,066 | 1091,237 | 996,05 | 20,89 | 1070,6148 | 1069,452 | -0,109 | 0,20 |
| 1600 | 1588,486 | 1572,787 | 994,38 | 18,11 | 1556,1313 | 1553,512 | -0,168 | 0,20 |

- Consistency of the calibrations

5.1 Consistency of UME

During the project, the turbine gas meter with serial number 031271 was tested two times in the pilot laboratory (UME). Obtained results are mentioned below.

Table 9

| flow rate (m3/h) | error of the meter 30.4.2007 | error of the meter 30.11.2007 | difference |
| --- | --- | --- | --- |
| (m3/h) | (%) | (%) | (%) |
| 80 | -3,830 | -1,709 | 2,121 |
| 160 | -1,463 | -0,720 | 0,743 |
| 400 | -0,421 | -0,130 | 0,291 |
| 650 | -0,290 | -0,125 | 0,165 |
| 1100 | -0,232 | -0,109 | 0,123 |
| 1600 | -0,143 | -0,168 | 0,025 |

Graph 1

- Bi-lateral Results

Transfer Standard (TS) with serial number 031271 was calibrated by two laboratories and the results are presented in the Table 10 and Graph 2.

Table 10

| flow rate (m3/h) | NMi error of the meter 01.06.2007 | UME error of the meter 30.04.2007 | UME error of the meter 30.11.2007 | Maximum error difference between NMi and UME first results | Maximum error difference between NMi and UME second results |
| --- | --- | --- | --- | --- | --- |
| (m3/h) | (%) | (%) | (%) | (%) | (%) |
| 80 | -2,09 | -3,830 | -1,709 | 1,74 | -0,381 |
| 160 | -0,68 | -1,463 | -0,720 | 0,783 | 0,04 |
| 400 | -0,22 | -0,421 | -0,130 | 0,201 | -0,09 |
| 650 | -0,20 | -0,290 | -0,125 | 0,09 | -0,075 |
| 1100 | -0,09 | -0,232 | -0,109 | 0,142 | 0,019 |
| 1600 | -0,12 | -0,143 | -0,168 | 0,023 | 0,048 |

Graph 2

Evaluation of degree of equivalence coefficient of bilateral comparison “Ei”

To evaluate the degree of eguivalence cofficient “Ei” for bilateral comparison, special  methoad  of M.G.Cox3  has been applied and the comparison equation is stated below.

Ei=Xlab−XrefUlab2+Uref2

Xref: Pilot country’s measurement error values (%)

Xlab: Other country’s measurement error values (%)

Uref: Pilot country’s measurement uncertainities

Ulab: Other country’s measurement uncertainities

Ei 1 Pass

Ei> 1 Fail

Table 11 - Bilateral Comparison result table for the First measurement

| Nominal Flowrate | NMi Xref | UME Xlab | NMi Uref | UME Ulab | Ei | Result |
| --- | --- | --- | --- | --- | --- | --- |
| (m3/h) | - | - | - | - | - | - |
| 80 | -2,09 | -3,83 | 0,19 | 0,20 | -6,31 | fail |
| 160 | -0,68 | -1,46 | 0,19 | 0,20 | -2,84 | fail |
| 400 | -0,22 | -0,42 | 0,19 | 0,20 | -0,73 | pass |
| 650 | -0,2 | -0,29 | 0,19 | 0,20 | -0,33 | pass |
| 1100 | -0,09 | -0,23 | 0,20 | 0,20 | -0,50 | pass |
| 1600 | -0,12 | -0,14 | 0,20 | 0,20 | -0,08 | pass |

Table 12 - Bilateral Comparison result table for the First measurement

| Nominal Flowrate | NMi Xref | UME Xlab | NMi Uref | UME Ulab | Ei | Result |
| --- | --- | --- | --- | --- | --- | --- |
| (m3/h) | - | - | - | - | - | - |
| 80 | -2,09 | -1,71 | 0,19 | 0,20 | 1,38 | pass |
| 160 | -0,68 | -0,72 | 0,19 | 0,20 | -0,14 | pass |
| 400 | -0,22 | -0,13 | 0,19 | 0,20 | 0,33 | pass |
| 650 | -0,2 | -0,13 | 0,19 | 0,20 | 0,27 | pass |
| 1100 | -0,09 | -0,11 | 0,20 | 0,20 | -0,07 | pass |
| 1600 | -0,12 | -0,17 | 0,20 | 0,20 | -0,17 | pass |

From the Table 11 and Table 12, the degree of equivalence coefficient of bilateral comparison “Ei” show that both comparisons are reasonable acceptable

- Summary and conclusion of EUROMET 1007 bilateral comparison

It has been seen that the bilateral comparison results are satisfactory for Turkey (UME) for the flow rates of 1600 m3/h, 1100 m3/h, 650 m3/h, 400 m3/h and 160 m3/h.

High error difference can be seen for the minimum measurement rate, 80 m3/h, between two measurements of UME. The frequency converter may be the reason for inconsistency at low flow rates.

- Comparison of EUROMET 862 and EUROMET 1007 measurement results

  - NMI - UME comparison for EUROMET 862

Table 13 was obtained from EUROMET 862 Final report Draft –A

Table 13

| flow rate (m3/h) | NMi error of the meter EUROMET 862 | UME error of the meter EUROMET 862 | Maximum error difference between NMi and UME for EUROMET 862 |
| --- | --- | --- | --- |
| (m3/h) | (%) | (%) | (%) |
| 100 | 1,18 | 0,96 | 0,22 |
| 200 | 0,45 | 0,73 | -0,28 |
| 300 | 0,26 | 0,45 | -0,19 |
| 400 | 0,32 | 0,62 | -0,3 |
| 500 | 0,37 | 0,68 | -0,31 |
| 650 | 0,46 | 0,72 | -0,26 |
| 801 | 0,43 | 0,68 | -0,25 |
| 1000 | 0,26 | 0,58 | -0,32 |

  - NMI - UME Eror difference comparison for EUROMET 862 and EUROMET 1007

From the Table 10 and Table 11, Table 13 and Graph 3 can be obtained.

Table 12

| flow rate (m3/h) | Error differences between NMi and UME For EUROMET 862 | Error differences between NMi and UME For EUROMET 1007 |
| --- | --- | --- |
| (m3/h) | (%) | (%) |
| 80 | 0,220 | -0,381 |
| 160 | -0,280 | 0,040 |
| 400 | -0,300 | -0,090 |
| 650 | -0,260 | -0,075 |
| 1100 | -0,320 | 0,019 |
| 1600 | - | 0,048 |

Graph 3

  - NMI consistency

The gas meter calibration results consistency for UME Master meter 031265 calibrated in 2004 and 2007; Table 13 and Graph 4 shows the consistency of both measurements.

Table 13

| flow rate (m3/h) | error of the meter 27.4.2004 | error of the meter 30.11.2007 | difference |
| --- | --- | --- | --- |
| (m3/h) | (%) | (%) | (%) |
| 80 | -0,28 | -0,50 | 0,22 |
| 160 | -0,04 | -0,12 | 0,08 |
| 400 | -0,18 | -0,17 | -0,01 |
| 650 | -0,20 | -0,19 | -0,01 |
| 1100 | -0,13 | -0,09 | -0,04 |
| 1600 | -0,12 | -0,07 | -0,05 |

Graph 4

Graph 2

  - UME Installation effects Investigation and consistency

Table 14

| flow rate | Error Between UME reference standard and UME-TS (Reference standard was connected with TS in serial) EURMET 1007 30.11.2007 | Error Between UME reference standard and UME-TS (Reference standard was put on place) EURMET 1007 30.11.2007 | Error Between NMi reference standard and UME-TS EURMET 1007 01.06.2007 |
| --- | --- | --- | --- |
| (m3/h) | (%) | (%) | (%) |
| 80 | -1,875 | -1,709 | -2,09 |
| 160 | -0,772 | -0,72 | -0,68 |
| 400 | -0,138 | -0,13 | -0,22 |
| 650 | -0,108 | -0,125 | -0,20 |
| 1100 | -0,072 | -0,109 | -0,09 |
| 1600 | -0,123 | -0,168 | -0,12 |

Graph 5

- Summary and conclusion of comparison between EUROMET 862  and EUROMET 1007  comparison results:

It can be seen from the Comparison Graph 3 “COMPARISON OF EUROMET 862 and EUROMET 1007”, there is are big differences (about 0,2%, 0,3%) between two comparisons.

Those differences might be related with;

- Master meter “K” factor

To investigate the UME master meter “K” factor, we sent the UME master meter to NMi for re-calibration, at the same time with the Transfer Standard for bi-lateral comparison.

The UME master meter 031265 calibrations at NMI for the second time showed that the results of two calibrations by NMI in 2004 and 2007 are consistent (Table 13 and Graph 4). And Graph 4 shows that there is no need for correction of UME reference K-factor used on the results of Euromet project 862.

- UME overall system performance, during the EUROMET 862 inter-comparison.

- Investigation of UME Medium Flow standards’ Installation affects; Installation effects of UME Piping system were investigated during the Second tour measurement of EUROMET 1007 at UME. First, UME reference meter and UME TS was connected in serial and data was taken, then UME reference meter was put on place than data was taken. And than Error and repeatability analyses were done. As seen from   Table 14 and Graph 5, there is no reasonable difference between them. This means that UME Installation effects of piping system did not affect the EUROMET project 862.

- Investigation of Flow computer: The same flow computer was used for EUROMET 862 and EUROMET 1007 comparison measurements without any changes.

- Investigation of measurants: Measurements were carried out by the same researchers in both comparisons.

- Air conditioning unites: There were some problems with the air conditioning unit during the measurements of EUROMET 862 inter comparison. But this may not be so important if the temperature, pressure and humidity sensors are working correctly.

- Investigation of Temperature and pressure sensors: After EUROMET 862 intercomparison measurements, we realised that from time to time those sensors do not work properly. Sometimes they give, sometimes wrong signals because of some noise problems due to the high magnetic field induced by suction fans. Therefore they are always under our control with the other calibrated reference standards. We have ordered new sensors to replace the existing sensors.

m3/h: Flowrate at working conditions

Sm3/h: Flowrate at standard conditions, 20 0C and 1013,25 mbar

m3/h: Flowrate at working conditions

Sm3/h: Flowrate at standard conditions, 20 0C and 1013,25 mbar

m3/h: Flowrate at working conditions

Sm3/h: Flowrate at standard conditions, 20 0C and 1013,25 mbar

3) Cox M.G., Metrologia, 2002, 39, 589-595